Low-Cost Visible Light Spectral Imaging

Daniel W. Dichter

Independent Researcher

Cambridge, Massachusetts, U.S.A.

daniel.w.dichter@gmail.com

Submitted: August 12, 2021

Spectral imaging is an emerging technology for measuring spectral power distributions (SPDs) of electromagnetic radiation over a two-dimensional spatial domain. Within the visible light wavelength domain, spectral imaging measures light and color with greater accuracy than digital photography. However, high cost limits its accessibility. Accordingly, a low-cost method was developed using commercially-available hardware – primarily a DSLR camera and a set of narrow bandpass filters. The quantity of filters was minimized to a total of seven, set by the dimensionality of SPDs, the spectral sensitivities of eyes and cameras, and commercial availability. Camera spectral sensitivity was measured using this same filter set, a color chart, a spectrophotometer, and noon daylight modeled as CIE D65. The RAW photo format was used to access unprocessed sensor data. Independent SPD measurements from each color channel were fused as a sensitivity-weighted average for efficient and continuous interpolation between color channels with a high signal-to-noise ratio. Images were reconstructed from SPDs with standard observer functions. The method was demonstrated with a Canon 650D DSLR camera, a set of Thorlabs one-inch narrow bandpass filters, an X-Rite ColorChecker chart, and a Spectro 1 spectrophotometer. Accuracy was validated by quantitative comparison against ground truth SPD measurements and qualitative assessment of reconstructed images. The total filter cost was $715, plus $405 to measure camera sensitivity.

# Introduction

## Background

Color perception results from the interaction of an illuminant, optionally a reflective surface, and an observer. Humans are trichromats, having three-dimensional color perception, over the visible light wavelength spectrum of 400 to 700 nm. Human color perception is significantly rank-deficient even within its wavelength limits, as shown in Sections 2.1 and 2.2.

Whereas digital photography produces a three-dimensional RGB (red, green, blue) color image, spectral imaging produces a *k*-dimensional hyperspectral datacube (HSDC), typically with 10 ≤ *k* ≤ 60; wavelength channels replace color channels. HSDCs are informationally complete and observer independent, providing a more accurate representation of color, and enabling greater optionality in subsequent image processing.

Spectral scanning is the simplest of several methods for generating HSDCs, wherein a 1-channel intensity image is captured at a series of regularly-spaced, narrow, non-overlapping wavelength bands. These images are subsequently “stacked” to form the final HSDC. Though spatially high-resolution relative to other methods, spectral scanning experiences “spectral smearing” if the subjects or sensor move during the image capture process, due to the characteristically large temporal distribution of the images. By comparison, so-called “single-shot” methods use various diffraction techniques to capture the entire HSDC simultaneously, but at lower spatial resolution.

A survey of commercially-available hyperspectral cameras[[1]](#footnote-1) showed a price range of $20,000-25,000, whereas consumer cameras such as DSLRs and MILCs cost 1-10% of this amount.

## Related Work

Consumer cameras are not capable of spectral imaging as-is, but various researchers have demonstrated this capability nonetheless with various hardware and software methods. Such methods can generally be classified as multi-shot spectral scanning or single-shot diffraction grating.

Cosentino [6] used 12 bandpass filters and a modified camera to produce HSDCs, comparing results favorably to a commercial hyperspectral camera. Berns et al. [3] used a large-format camera and an optimized set of filters, performing singular value decomposition on a dataset of 2,500 reflectance spectra to reduce the quantity of filters to a total of six.

Baek et al. [2] developed a single-shot method using a custom prism objective. Combined spatial and wavelength information was separated by detection of “spectral cues” present “only around object edges”. Habel et al. [7] similarly developed a single-shot diffraction grating method, with reconstructed images limited to 120 x 120 pixels.

Oh et al. [10] developed a novel non-diffraction single-shot method using three different synchronized digital cameras, exploiting small differences in their sensitivities. An image registration process was used to align images between cameras using planarity. Principal component analysis was performed on a database of 1,257 Munsell reflectance spectra to create a low-dimension vector space for describing SPDs as linear combinations.

The method described in this paper is believed to be novel in its low cost, high spatial resolution, use of unmodified commercially-available hardware, and independence from training data.

# Methods

## Dimensionality of Reflectance Spectra

Parkkinen et al. [11] measured the reflectance spectra of 1,257 standard Munsell color swatches over 400 to 700 nm. Noting that “the components of a color spectrum are highly correlated”, the authors performed principal component analysis on these spectra, producing a set of reflectance eigenvectors. It was found that these spectra could be reconstructed as linear combinations of eight or fewer eigenvectors. This result indicates that reflectance spectra are up to eight-dimensional over the visible light spectrum, with a characteristic wavelength resolution of approximately (700 nm - 400 nm) / 8 = 37.5 nm.

## Curve Reconstruction

Sampling an SPD by spectral scanning with narrow bandpass filters produces a sparse sample; the value of the SPD is measured only at certain regular intervals. Between these, the value of the SPD is not measured, but can be approximated due to the limited dimensionality of SPDs. This process was modeled as a curve reconstruction problem, i.e. given a sparse sampling of an unknown curve, reconstruct the curve by means of an appropriate interpolation scheme such that the reconstructed curve matches a theoretical measured curve. This concept is shown in Figure 1.



**Figure 1**: Curve reconstruction for SPD of ColorChecker swatch #2 under D65 illuminant for varying sample quantity. Samples are denoted with vertical lines, with cubic interpolation between samples.

To determine the optimal filter quantity and interpolation scheme, two reflectance datasets were considered: the 8 eigenvectors developed by Parkkinen et al. [11] with 5 nm resolution, and the 24 reflectance spectra measured from an X-Rite ColorChecker Classic. [13]

ColorChecker reflectance spectra were measured using a Spectro 1 spectrophotometer with a domain of 400 to 700 nm and a resolution of 10 nm. Three scans were performed and averaged for each swatch, with near-perfect repeatability across trials. These spectra were validated by calculating the CIEDE2000 color difference (denoted ΔE00) against the manufacturer's published Lab color values under CIE D50 illuminant with a 2° observer. [13] The error for all 24 swatches was 1.39 ± 1.06 ΔE00, showing good agreement. Research has shown significant variance in ColorChecker reflectance spectra, with standard deviations of 0.1-9.1 percentage points over 400-700 nm, [1] though the manufacturer does not publish spectra or tolerances.

Per Sections 2.1 and 2.3, both 30 and 40 nm resolution were considered. The reconstruction domain was taken as 420-660 nm per Section 2.3. The filters were assumed to be of sufficiently narrow bandwidth that they measure at their exact center wavelength (CWL). Linear, spline, and cubic interpolation were considered. The quality of the reconstruction was calculated as the residual sum of squares (RSS) for all curves, multiplied by the wavelength resolution, divided by the quantity of samples in the dataset. This calculation was performed for both reflectance spectra and simulated SPDs, with the latter modeled as the element-wise product of reflectance and the D65 illuminant scaled to the range of 0-1. This normalized error metric allowed direct comparison between datasets.

Several observations were made. 30 nm resolution outperforms 40 nm resolution in all cases. With 40 nm resolution, cubic interpolation is optimal. With 30 nm resolution, cubic interpolation is at least near-optimal. Optimized error at 40 nm is 1-3x larger than at 30 nm, but both are small in an absolute sense. Reflectance (not illuminant) dominates the quality of the curve reconstruction. Thus, 40 nm resolution and cubic interpolation were chosen for curve reconstruction.

## Filter Set Design

Manufacturers of narrow bandpass filters include Thorlabs, Edmund Optics, and MidOpt. Within the visible light spectrum, CWL is generally discretized as whole-number multiples of 10 nm, i.e. 400, 410, 420, ..., 700 nm.

The centroid of the visible light spectrum may be defined at the wavelength corresponding to 50% on the cumulative density function (CDF) of the sum of the CIE 2° tristimulus observer functions . [5] Rounding to the nearest 10 nm per commercial availability, this centroid is 540 nm.

It can be shown that the wavelength range of 420 to 660 nm encompasses 97% of the area under the CIE 2° and 10° tristimulus observer functions, [5] thus:

|  | (1) |
| --- | --- |

This range is also evenly divisible into 40 nm increments, and intersects the centroid of 540 nm.

The final specification is the full-width half max (FWHM), or bandwidth. As discussed in Section 2.2, reducing FWHM increases the wavelength accuracy of measurements to the CWL. The reduction in overall transmission associated with a low FWHM was compensated by increasing the exposure, as discussed in Section 2.6.

The set of filters chosen are shown in Figure 2 and Figure 7, and have the following properties:

Quantity 7

Manufacturer Thorlabs

CWL 420 to 660 nm

CWL Spacing 40 nm

FWHM 10 nm

Part Numbers FB420-10, FB460-10, etc.

Diameter One inch

Cost, Total $686 (+$29 case)



**Figure 2:** Thorlabs narrow bandpass filter transmission spectra, per manufacturer’s datasheets. [12]

## Camera Response Linearity

RAW values reported from cameras may be thought of as photonic measurements. A simple under-exposure experiment found that pure black (i.e. zero photons) corresponds to RAW = 2,048. In principle, over-exposure and saturation occur at the maximum value permitted by the bit depth: 214 = 16,384. In practice, saturation was observed at values ranging from 12,000-16,384. Between these limits, response linearity was verified by photographing a ColorChecker chart under noon daylight while independently varying shutter duration and ISO. The trichromate mode of the #19 white swatch was calculated for each photo as the measurement of interest. Mode was chosen for its robustness against hot and dead pixels.

These measurements are shown in Figure 3. For ideal linearity, the RAW value is proportional to the product of shutter duration and ISO, all else equal. Accordingly, an idealized RAW value was calculated for each reported RAW value, by linearly scaling the product of shutter duration and ISO to the RAW value range of 2,048-12,000. This relationship was expressed as a set of transfer functions, shown in Figure 4. The shape of these transfer functions indicates that the sensor has a linear response up to saturation. Similarly, varying ISO with a constant shutter duration exhibited near-perfect linearity (f/22, ISO 100-12,800, 1/2,000 sec).



**Figure 3:** Reported RAW values for #19 white ColorChecker swatch in noon daylight without filters; f/22, ISO 100, 1/4,000-1/30 sec.



**Figure 4:** Photonic transfer functions for data in Figure 3.

## Camera Spectral Sensitivity

Camera spectral sensitivity was calculated by photographing a ColorChecker chart through each filter under cloudless noon daylight modeled as D65. Noon daylight was chosen for its roughly neutral SPD, availability of standard data, and accessibility. Images were captured on 2021-02-21 at 12:22 pm in an open field in Cambridge, Massachusetts.

For an ideal linear response as shown in Section 2.4, sensitivity may be described generally as:

|  | (2) |
| --- | --- |

VM was calculated as the mode of the RAW values inside a square inset slightly from the swatch perimeter. VA was calculated as:

|  | (3) |
| --- | --- |

With:

λ Wavelength

I Scene illuminant

R Swatch reflectance

T Filter transmission

⊙ Element-wise multiplication

Sensitivity for all color channels and wavelengths was calculated from each swatch, as shown in Figure 5. These per-swatch sensitivity curves were fused as a weighted average using VM to weight calculated sensitivities in proportion to their signal-to-noise ratio.

The result is a set of three sensitivity curves, . As shown in Figure 6, this result is consistent with values from literature for similar Canon cameras.



**Figure 5:** Non-dimensional trichromate spectral sensitivity of Canon 650D camera from ColorChecker under D65 illuminant. Faint colored lines correspond to individual swatches; solid black lines correspond to weighted averages for all swatches.

## 

**Figure 6:** Comparison of sensitivities for various Canon cameras. [8] Curves were normalized so that the maximum sensitivity for each camera is equal to unity.

## Photographic Aspects

A prime (i.e. non-zoom) lens, tripod, quick release plate, and remote shutter were used to minimize movement of the camera while capturing the photo stack. Because filter cost is proportional to filter area, the choice of lens is the primary means of cost reduction. The measurement of interest is the diameter of the objective at the outer end of the lens, as it sets the minimum filter diameter. For consumer Canon EF lenses, this ranges from approximately 20 to 60 mm. The theoretical filter cost at these extremes differs by nearly an order of magnitude; 602 / 202 = 9. The lens selected has an outer objective diameter of 20 mm, a focal length of 40 mm, and an aperture of f/2.8. It is shown in Figure 7.

The filters discussed in Section 2.3 are unthreaded, and used by resting them against the camera lens by hand, transferring minimal force and maintaining alignment of the photo stack. The glass optics of both the lens and filters are recessed from their cases, preventing abrasion. The filters were stored in a case in order of ascending wavelength, and cycled through in sequence manually. Depending on the nature and luminance of the scene, typical camera settings were f/2.8-10, ISO 100-400, and 1/4 - 1/500 second. Bright exposures were required to compensate for the low transmission of the filters. These settings were adjusted to “expose to the right”, i.e. fully utilize the available set of values without saturating the upper limit, thereby maximizing the signal-to-noise ratio in the measured values.

The open-source dcraw utility was used to extract RAW values from the .CR2 file format. [4] The modifier string -D -4 -j -t 0 was used to specify that the extracted RAW values were unprocessed sensor measurements:

-D No value scaling

-4 Linear 16-bit

-j No stretching or rotating pixels

-t 0 No image rotation

The RAW sensor measurements were then demosaiced into R, G, B color channels with a Bayer filter pattern of rggb.



**Figure 7:** Canon 650D camera with 40 mm prime EF lens and set of Thorlabs narrow bandpass filters.

## SPDs from RAW Photos

By rearranging the general expression of sensitivity in Section 2.5:

|  | (4) |
| --- | --- |

For a linear sensor response, , with P denoting the RAW photo value. By inspection, . Using proportionality and non-dimensionality, this is rearranged and substituted as:

|  | (5) |
| --- | --- |

The denominator ST is most accurately and generally expressed as a dot product over the wavelength domain rather than a scalar product at the CWL, i.e. S(λ) · T(λ) rather than S(λ=CWL) \* T(λ=CWL).

At an arbitrary pixel and CWL, each color channel produces an independent SPD measurement, collectively given as:

|  | (6) |
| --- | --- |

These measurements are theoretically equal, but in practice differ due to various sources of error. They are fused as a sensitivity-weighted average:

|  | (7) |
| --- | --- |

Calculating as a sensitivity-weighted average continuously interpolates between color channels as a function of wavelength. Since SR+SG+SB > 0 for all λ, divide-by-zero is precluded. This formulation was generalized spatially by using matrices in place of scalars for P, neglecting vignetting, dark-frame effects, and other spatially-related sources of error.

The HSDC was first calculated in a sparse fashion, only at the CWLs of the filters. The full (i.e. non-sparse) HSDC was then calculated by cubic interpolation in the wavelength domain between the filter CWLs per Section 2.2.

# Results

## Validation

The method was validated by comparing ColorChecker SPDs and ΔE00 between camera vs. spectrophotometer, with the latter taken as ground truth and modeled as R(λ) ⊙ I(λ). Because the two sets of curves, SPDcamera and SPDspectrophotometer, are non-dimensional, their ranges are aligned with a single scalar gain for comparative purposes, such that the sum of the residuals is zero. This scalar gain α was found by the bisection method such that:

|  | (8) |
| --- | --- |

Standard equations were used to derive XYZ, Lab, and RGB colors from SPDs, with D65 as both the scaling factor and white point. [9]

Results are shown in Figure 8, Figure 9, and Figure 10. The color error for all swatches is 2.44 ± 1.65 ΔE00, with more than 90% of swatches within 1.75 ± 0.96 ΔE00. The primary outlier is #19 white, which shows a good match in terms of normalized distribution, but a poor match in magnitude and thus luminance. This is best explained as an artifact of glare or other lighting non-uniformities, amplified by the high reflectance of the color white. The SPDs of the matte black background showed a spatial luminance variance that matched the SPD magnitude error.

In the more general case of an unknown illuminant, colors are dimensionalized by linear scaling in the luminance (or intensity) dimension to the limits of the color space. A common strategy is to allow a small amount of saturation at the limits of the color space, thereby avoiding hot and dead pixels. Reconstructed images in this paper are scaled such that the bottom and top 1% of pixels are saturated, with black and white defined as 5% and 95% (i.e. RGB [13 13 13], [242 242 242]) respectively.



**Figure 8:** Distribution of color errors shown in Figure 10.

## 

**Figure 9:** Reconstructed image of ColorChecker chart; f/2.8, ISO 100, 1/400 sec.



**Figure 10:** Comparison of ColorChecker SPDs and colors as measured by spectrophotometer (solid, left) and camera (dashed, right). A single scalar gain is applied to align the two sets of curves. Plots are scaled 400 to 700 nm along x, and non-dimensional along y.

## Reconstructed Images

Several scenes are shown in Figure 11 with subjects and illuminants that represent typical color perception. Shown for each scene is the reconstructed image (left column) and a sparse sampling of the HSDC (right column). Each HSDC was sampled at an evenly-distributed 10 x 7 square mesh with 70 nodes total, showing a characteristic set of SPDs that are colored according the reconstructed image.

The still life (first row) exhibits blue-red contrast apparent in the bimodal distribution of the SPDs. A relatively high f-stop was needed to keep the scene in focus, which required increasing both ISO and shutter duration. The lower signal-to-noise ratio inherent with higher ISO can be seen in the faint high-frequency noise pattern. The camera settings were f/10, ISO 400, 1/4 sec.

The landscape (second row) exhibits both color and luminance contrast between sky and ground. Minor “spectral smearing” can be seen around the edges of the foliage due to natural movement. The sky exhibits a characteristic D65-like distribution. Higher overall luminance enabled lower ISO and higher shutter duration, reducing noise. The camera settings were f/2.8, ISO 200, 1/200 sec.

The industrial scene (third row) exhibits correlation between SPDs that is consistent with the yellow-red light characteristic of sunsets. The camera settings were f/4, ISO 200, 1/60 sec.







**Figure 11:** Reconstructed RGB images and characteristic sets of SPDs for several scenes. Plots are scaled 420 to 660 nm along x, and non-dimensional along y. Filter CWLs are denoted by vertical lines.

# Discussion

This paper demonstrates a method for spatially high-resolution visible light spectral imaging at less than 10% the cost of a commercial hyperspectral camera, using commercially-available hardware. Results are validated quantitatively and qualitatively against ground truth. Key aspects of novelty include characterization of SPD dimensionality, curve reconstruction from sparse samples, and fusion of trichromate measurements. This method significantly improves access to the field of spectral imaging, and enables further research into the field of spectral image processing.

# Acknowledgments

The author gratefully recognizes his friends, family, and colleagues for their generous and thoughtful support in shaping this work – in particular: Eric Kirchner, Trey Henderson, Bronek Dichter, Michał Dichter, and Evan Molinari.

# References

1. BabelColor. *The ColorChecker Pages*. Accessed 2021-08-01. https://www.babelcolor.com
2. Baek, S.-H., Kim, I., Gutierrez, D., & Kim, M. H. (2017). *Compact Single-Shot Hyperspectral Imaging Using a Prism*. ACM Transactions on Graphics, 36(6), 1–12.
3. Berns, R., Taplin, L., Nezamabadi, M., Mohammadi, M., & Zhao, Y. (2005). *Spectral imaging using a commercial color-filter array digital camera*.
4. Coffin, D. *Decoding raw digital photos in Linux*. (n.d.). https://www.dechifro.org/dcraw/.
5. Colour & Vision Research Laboratory (n.d.). *CVRL Main*. http://cvrl.ioo.ucl.ac.uk/index.htm.
6. Cosentino, A. (2015). *Multispectral Imaging System Using 12 Interference Filters for Mapping Pigments*. Conservar Património, 21, 25–38. https://doi.org/10.14568/cp2015005
7. Habel, R., Kudenov, M., & Wimmer, M. (2012). *Practical Spectral Photography*. Computer Graphics Forum, 31(2pt2), 449–458. https://doi.org/10.1111/j.1467-8659.2012.03024.x
8. Jiang, J., Liu, D., Gu, J., & Susstrunk, S. (2013). *What is the Space of Spectral Sensitivity Functions for Digital Color Cameras?* 2013 IEEE Workshop on Applications of Computer Vision (WACV). https://doi.org/10.1109/wacv.2013.6475015
9. Lindbloom, B. J. (n.d.). *Spectral Computation of XYZ*. http://www.brucelindbloom.com.
10. Oh, S. W., Brown, M. S., Pollefeys, M., & Kim, S. J. (2016). *Do It Yourself Hyperspectral Imaging with Everyday Digital Cameras*. 2016 IEEE Conference on Computer Vision and Pattern Recognition (CVPR). https://doi.org/10.1109/cvpr.2016.270
11. Parkkinen, J. P. S., Hallikainen, J., & Jaaskelainen, T. (1989, February 1). *Characteristic Spectra of Munsell Colors*. JOSA A.
12. Thorlabs. *UV/VIS Bandpass & Laser Line Filters: 340 - 694.3 nm Center Wavelength*. Accessed 2020-10-19. https://www.thorlabs.com.
13. X-Rite. *Colorimetric values for ColorChecker family of targets*. 2019-04-11. https://xritephoto.com.

1. PixelTeq SpectroCam-VIS, BaySpec GoldenEye, Resonon Pika L [↑](#footnote-ref-1)